ST-9 Linearity Analysis

The transfer curve matches the equation
{ ADU} = -3.59982546367x^2 + 1198.70146x - 396.0608
where x is an adjusted exposure time that we can assume has some linear relation to flux, i.e.,
{ Flux } = { m}{x }+b
where x is still the adjusted exposure time.

Let ADU' be the number of ADU that would have been reported if the camera was perfectly linear. Further, let's assume that {partial { ADU' } } over { partial {Flux }} = {partial { ADU } } over { partial {Flux }} for small values of flux. If Flux is small, then x should be small and x2 will be smaller still. That means that the slope of both curves should be approximately 1198.7x.

If we eliminate bias corrections, then set ADU' = 1198.7x, this means that
x = { ADU' }over{ 1198.7 } , which means (from the first equation):
{ ADU } = -3.59982546367 left ( {ADU'}over{1198.7} right )^2 + { 1198.70146 {ADU'}}over{1198.70146 }
ADU = -2.505298 times 10^{ -6} {ADU' }^2 + { ADU'}

Now, solve this for ADU' in terms of ADU:
{ ADU' }= { -1 +-{sqrt{1+4{ADU}}} }over{ 2 times (-2.505298 times 10^{-6}) }
Yuk. Shift gears.

If we plot the curve of x vs reported ADU, we get a curve that matches the equation:
x = 3.78120034622418 times 10^{ -9 } { ADU }^2 + 0.000791910055837{ ADU } + 0.5073
then substitute our earlier relation between ADU' and x, we see:
{ADU'}over{1198.7} = 3.78120034622418 times 10^{ -9 } { ADU }^2 + 0.000791910055837{ ADU } + 0.5073
which can be simplified to:
{ ADU' }= 4.7748 times 10^{ -6}ADU^2 + ADU + { constant}

We arbitrarily set the constant to zero, which means that a measured ADU needs to be adjusted by adding 4.7748 times 10^{ -6 } { ADU }^2 to the ADU to turn it into an equivalent ADU'.

With no bias adjustment:
{ Adjustment }= A ( x - 32768 )^2 + 3224
with
A={ -3224 }over{ 32768^2 }
This adjustment is subtracted from the reported ADU to get a "true" ADU. Note that this will also change the system gain by about 20%.

Now we add in the bias adjustment. Bias level on the ST-9 is about 105, so we want the adjustment to be zero at ADU values of 105 and 65535. Midpoint between these two is 32820. So,
{ Adjustment }= A ( x - 32820 )^2 + 3224 , so that means that
A={ -3224 }over{ 32715^2 }

Do this yet again.
y=-3.59982546367124x^2 + 1198.70146236367x - 396.060839
First, drop the constant term, because we want this equation to pass through the origin.
And also set the slope near zero to be 1.0, by substituting x' = x/1198.70146236367:
y={-3.59982546367124{x'}^2}over{1198.70146236367^2} + x
So,
y = x - 2.5052979 times 10^{ -6}x^2
The error term alone is
{ error }= -2.5052979 times 10^{ -6}x^2 so that for some flux f and some system gain G0,

{ ADU } = -2.5052979 times 10^{ -6} left ( {f}over{G_0} right ) ^2 + left ( {f}over{G_0}right )
We assert that there is some new G1 and f0 for which
-2.5052979times10^{ -6}left( {f}over{G_0}right )^2 = { K}left( {f}over{G_1}right )+A left( {f-f_0}over{G_1} right )^2 + Q

Do this yet again.
{ ADU } = -2.5052979 times 10^{ -6} left ( {f}over{G_0} right ) ^2 + left ( {f}over{G_0}right )
Nominal "full-well" flux fmax is the largest flux of interest. It produces an ADU of 65535. Substitute z={ f }over{ G_0 } and define zmax as the value of z when f = fmax. Then,
65535 = -2.5052979times10^{ -6}z_max^2 + z_max
This means that
z_max={ {-1} +- {sqrt{1-(4)(-2.5052979)times10^{-6}(-65535)}} } over { (2) (-2.5052979) times 10^{-6} }=82644.9

{ ADU } = left ( -2.5052979 times 10^{ -6} left ( {f}over{G_0} right ) ^2 + left ( {f}over{G_0}right ) right ) + { correction(ADU) } = { f }over{ G_1 }
Further, we want:
correction( 0 ) = 0 and correction( 65535) = 0 and the correction function is of the form
correction( ADU) = %alpha { (ADU)}^2 + %beta { (ADU)} which means that 65535 times %alpha+%beta = 0 . Further, we know that the f2 term must zero out for all f, so that α = 2.5052979x10-6, and β = (-65535)( 2.5052979x10-6) = -0.1642. This also means that { f }over{ G_0 }+{ (-0.1642)f }over{ G_0 }={ f }over{ G_1 } , or
{ G_0 }over{ .8358 }=G_1 .

Hence, the correction is
correction( ADU )= 2.5052979times10^{ -6 }( {ADU}^2 )+( -0.1642 )( {ADU} )